## Slide 1
ME23312 ENGINEERING THERMODYNAMICS
Unit III – Properties of Pure Substance · Presentation 8 of 10
Determination of the Dryness Fraction – Steam Calorimeters
Sampling · barrel calorimeter · separating calorimeter · throttling calorimeter · combined separating and throttling calorimeter · comparison
Dr. Anantha Raman L., M.E., Ph.D. · Associate Professor, Department of Mechanical Engineering, Rajalakshmi Institute of Technology, Chennai

## Slide 2
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
2
In this presentation
Why a calorimeter?
Barrel (bucket) calorimeter
Separating calorimeter
Throttling calorimeter
When the throttling calorimeter works
Combined separating and throttling calorimeter
Comparison of the calorimeters

## Slide 3
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
3
Why a calorimeter?
The dryness fraction of steam leaving a boiler or flowing in a pipe cannot be measured directly
A steam calorimeter brings a sample to a state whose properties can be measured and applies the energy balance
The sample is taken through a perforated tube inserted across the pipe so that it is representative of the flow
Instruments: barrel, separating, throttling, combined, electrical

## Slide 4
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
4
1
Barrel calorimeter

## Slide 5
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
5
Barrel (bucket) calorimeter
ms [hf + x hfg – cw t2] = (mw + me) cw (t2 – t1)
Wet steam ms condensed in cold water mw (initial t1) in an insulated barrel; final temperature t2
Heat lost by the steam = heat gained by the water and the barrel (water equivalent me)
Simple but inaccurate – heat losses and a small temperature rise

## Slide 6
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
6
2
Separating calorimeter

## Slide 7
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
7
Separating calorimeter
Baffles throw the heavier water droplets out of the stream; the water collects below and the drier steam passes on to be condensed

## Slide 8
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
8
Separating calorimeter
x = mg/(mf + mg)
Sudden changes of direction throw out the water; its mass mf is read from a gauge glass; the drier steam is condensed and weighed, mg
Separation is never complete – fine moisture is carried over, so it over-estimates x
The only instrument for very wet steam; first stage of the combined calorimeter

## Slide 9
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
9
3
Throttling calorimeter

## Slide 10
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
10
Throttling calorimeter
The sample is throttled from the pipe pressure p1 to a low pressure p2; p2 and T2 are measured in the insulated chamber

## Slide 11
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
11
Throttling calorimeter
hf1 + x1 hfg1 = h2(p2, T2) ⇒ x1 = (h2 – hf1)/hfg1
Throttling is a constant-enthalpy process (Unit I)
If the throttled steam is superheated, its state is fixed by p2 and T2 and h2 is read from the superheat table
On the Mollier chart: a horizontal line from the wet region at p1 into the superheated region at p2

## Slide 12
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
12
When the throttling calorimeter works
Accurate – but only if the throttled steam is superheated by at least about 5 K
Limits it to fairly dry steam: x above roughly 0.95 for the usual pressure ratios
For wetter steam the two instruments are combined

## Slide 13
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
13
4
Combined separating and throttling calorimeter

## Slide 14
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
14
Combined separating and throttling calorimeter
The separator removes most of the moisture; the partially dried steam is then throttled into the superheated region

## Slide 15
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
15
Combined calorimeter
x2 = (h3 – hf1)/hfg1 x = x2 m2/(m1 + m2) = x1 x2
Separator collects water m1; the partially dried steam (x2, mass m2) is throttled to p3, T3
hf1, hfg1 at the pipe pressure p1 – unchanged through the separator
Dry vapour in the sample = x2 m2; total mass = m1 + m2; x1 = m2/(m1 + m2) is the separator reading

## Slide 16
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
16
Comparison of the calorimeters
Calorimeter | Principle | Suitable for | Limitation
Barrel | Energy balance with cold water | Rough checks | Inaccurate – heat losses, small ΔT
Separating | Mechanical separation of moisture | Very wet steam | Over-estimates x – carry-over of fine moisture
Throttling | h = constant into the superheated region | Fairly dry steam, x > 0.95 | Throttled steam must be superheated (≥ 5 K)
Combined | Separation, then throttling | Any wetness | Two instruments; x = x1 x2

## Slide 17
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
17
Summary
Dryness fraction is found from a representative sample with a calorimeter
Barrel: ms [hf + x hfg – cwt2] = (mw + me) cw (t2 – t1); inaccurate
Separating: x = mg/(mf + mg); over-estimates; best for very wet steam
Throttling: x1 = (h2 – hf1)/hfg1; needs superheat after throttling
Combined: x = x2 m2/(m1 + m2) = x1 x2
On the Mollier chart throttling is a horizontal line

## Slide 18
Next: Non-flow processes of steam
Work and heat transfer for constant-volume, constant-pressure, isothermal, hyperbolic, isentropic and polytropic processes
Numerical problems: Worked Examples 3.5 – 3.6 in the lecture notes
ME23312 Engineering Thermodynamics · Unit III · Presentation 8
18